## Solution std set B conc
Analyte | C (ppm) | U (ppm) | C (ppb) | U (ppb) |  | 4B (Multi) | 4B | 5B | B (As) | B (P) | B (Si) | B (Zr) | HNO3
Ag | 20 | 0.1 | 20000 | 100 | Name | Custom 3A | 4A | 5A | 6A | P | Si | Zr | 2% HNO3
Al | 99.83 | 0.5 | 99830 | 500 | Prod. # | VHG-Z21-DOc1-ss-w000364 | VHG-ZUTAUSTIN902-100 | VHG-ZUTAUSTIN903-100 | VHG-ZASN100-100 | VHG-ZPN-1000-100 | VHG-TSIW-SL-100 | VHG-ZZRN20-DOC1-100 | ICP-MS lab
B | 99.87 | 0.5 | 99870 | 500 | Lot # | 10015975-2 | 10039720-35 | 10039720-33 | 10037125-236 | 1027188-52 | 1094765-12 | 10015975-5
Ba | 19.99 | 0.1 | 19990 | 100 | Matrix | 5% Nitric | 5% HNO3/Tr. HF/tr. Tartaric | 5% Nitric | 5% Nitric | 5% Nitric | H20/tr. F- | 5% Nitric
Bi | 20.01 | 0.1 | 20010 | 100 | Density (g/mL) | 1.012 | 1.016 | 1.081 | 1.0024 | 1.0058 | 1.0354 | 1.022 | 1.012
Cd | 20.01 | 0.1 | 20010 | 100 | Temp. (ºC) | 19.9 | 21.5 | 20.7 | Room temp. | Room temp. | Room temp. | 19.7 | Room temp.
Co | 20.01 | 0.1 | 20010 | 100
Cr | 20 | 0.1 | 20000 | 100
Cs | 19.98 | 0.1 | 19980 | 100
Cu | 20.02 | 0.1 | 20020 | 100
Fe | 99.98 | 0.5 | 99980 | 500
Li | 100.1 | 0.5 | 100100 | 500
Mn | 20 | 0.1 | 20000 | 100
Ni | 20.01 | 0.1 | 20010 | 100
Pb | 19.89 | 0.1 | 19890 | 100
Rb | 19.99 | 0.1 | 19990 | 100
Se | 20 | 0.1 | 20000 | 100
Sr | 99.95 | 0.5 | 99950 | 500
Th | 19.98 | 0.1 | 19980 | 100
Tl | 20.02 | 0.1 | 20020 | 100
U | 19.98 | 0.1 | 19980 | 100
V | 20.01 | 0.1 | 20010 | 100
Zn | 99.92 | 0.5 | 99920 | 500
Mo | 20 | 0.1 | 20000 | 100
Sb | 20.01 | 0.1 | 20010 | 100
Sn | 19.97 | 0.1 | 19970 | 100
Ti | 20.07 | 0.1 | 20070 | 100
Ca | 10000 | 50 | 10000000 | 50000
K | 2500 | 13 | 2500000 | 13000
Mg | 5000 | 25 | 5000000 | 25000
Na | 10000 | 50 | 10000000 | 50000
As | 10 | 0.1 | 10000 | 100
P | 999 | 3 | 999000 | 3000
Si | 9973 | 50 | 9973000 | 50000
Zr | 20 | 0.1 | 20000 | 100

## Solution std set A conc
Analyte | C (ppm) | U (ppm) | C (ppb) | U (ppb) |  | 4B (Multi) | 4B | 5B | B (As) | B (P) | B (Si) | B (Zr) | HNO3
Ag | 19.89 | 0.1 | 19890 | 100 | Name | Custom Std. | Custom 902SS Standard 4B | Custom 903SS Standard 5B | Single-Element Aqueous CRM | Single-Element Aqueous CRM | Single-Element Aqueous CRM | Single-Element Aqueous CRM | n/a
Al | 99.93 | 0.5 | 99930 | 500 | Prod. # | VHG-Z21-DOC1-W000364 | VHG-ZUTAUSTIN902SS-100 | VHG-ZUTAUSTIN903SS-101 | VHG-LASN-100 | VHG-PPN-100 | VHG-TSIW-100 | VHG-ZZRN20-DOC1-SS-100 | n/a
B | 100.1 | 0.5 | 100100 | 500 | Lot # | 10015975-4 | 10039720-36 | 10039720-34 | 1093758-4 | 1112439-6 | 1004987-44 | 10015975-7 | n/a
Ba | 20.13 | 0.1 | 20130 | 100 | Matrix | 5% HNO3 | 5% HNO3 /tr. HF /tr. Tartaric acid | 5% HNO3 | 2% HNO3 | 5% HNO3 | H2O/tr. F- | 5% HNO3 | 2% HNO3
Bi | 20.01 | 0.1 | 20010 | 100 | Density (g/mL) | 1.017 | 1.02 | 1.091 | 1.014 | 1.0018 | 1.0353 | 1.035 | 1.012
Cd | 20.09 | 0.1 | 20090 | 100 | Temp. (ºC) | 19.9 | 21.5 | 20.7 | Room temp. | Room temp. | Room temp. | 19.7 | Room temp.
Co | 20.21 | 0.1 | 20210 | 100
Cr | 20.12 | 0.1 | 20120 | 100
Cs | 20.05 | 0.1 | 20050 | 100
Cu | 20.24 | 0.1 | 20240 | 100
Fe | 99.96 | 0.5 | 99960 | 500
Li | 99.86 | 0.5 | 99860 | 500
Mn | 20.13 | 0.1 | 20130 | 100
Ni | 20 | 0.1 | 20000 | 100
Pb | 20.09 | 0.1 | 20090 | 100
Rb | 19.8 | 0.1 | 19800 | 100
Se | 20.13 | 0.1 | 20130 | 100
Sr | 99.84 | 0.5 | 99840 | 500
Th | 19.85 | 0.1 | 19850 | 100
Tl | 20.05 | 0.1 | 20050 | 100
U | 20.25 | 0.1 | 20250 | 100
V | 20.23 | 0.1 | 20230 | 100
Zn | 100.2 | 0.5 | 100200 | 500
Mo | 19.99 | 0.1 | 19990 | 100
Sb | 20.01 | 0.1 | 20010 | 100
Sn | 20 | 0.1 | 20000 | 100
Ti | 20 | 0.1 | 20000 | 100
Ca | 10000 | 50 | 10000000 | 50000
K | 2500 | 13 | 2500000 | 13000
Mg | 4999 | 25 | 4999000 | 25000
Na | 9999 | 50 | 9999000 | 50000
As | 100 | 0.5 | 100000 | 500
P | 1005 | 3 | 1005000 | 3000
Si | 10020 | 50 | 10020000 | 50000
Zr | 20.04 | 0.1 | 20040 | 100

## Calibration stds (sol set)
From semi-quant: | [7 empty] | Hypothetical dilutions:
Analytes | Conc in ppb | After 1000x dil |  | Conc in 4B (ppb) | Conc in 5B (ppb) |  |  | L1 | L2 | L3 | L4 | L5 | L6 | Reqd DF for L6 |  |  |  |  | Actual
Li | 90 | 0.09 |  | 99860 |  |  |  | 0 | 0.24915 | 2.4915 | 24.9152 | 249.1517 | 498.3034 | [5 empty] | Analytes | Conc in 4B | Conc in 5B
Na | 184240 | 184.24 |  |  | 9999000 |  |  | 0 | 2.4998 | 24.9975 | 249.975 | 2499.75 | 4999.5 | [5 empty] | Li | 100100
Mg | 172960 | 172.96 |  |  | 4999000 |  |  | 0 | 1.2498 | 12.4975 | 124.975 | 1249.75 | 2499.5 | [5 empty] | Na |  | 10000000
Al | 45120 | 45.12 |  | 99930 |  |  |  | 0 | 0.24933 | 2.4933 | 24.9326 | 249.3263 | 498.6527 | [5 empty] | Mg |  | 5000000
Ca | 376000 | 376 |  |  | 10000000 |  |  | 0 | 2.5 | 25 | 250 | 2500 | 5000 | 2000 |  |  |  |  | Al | 99830
Cr | 4136 | 4.136 |  | 20120 |  |  |  | 0 | 0.0502 | 0.502 | 5.02 | 50.1996 | 100.3992 | [5 empty] | Ca |  | 10000000
Mn | 3045.6 | 3.0456 |  | 20130 |  |  |  | 0 | 0.050225 | 0.50225 | 5.0225 | 50.2246 | 100.4491 | [5 empty] | Cr | 20000
Fe | 225600 | 225.6 |  | 99960 |  |  |  | 0 | 0.2494 | 2.494 | 24.9401 | 249.4012 | 498.8024 | [5 empty] | Mn | 20000
Ni | 902.4 | 0.9024 |  | 20000 |  |  |  | 0 | 0.0499 | 0.499 | 4.99 | 49.9002 | 99.8004 | [5 empty] | Fe | 99980
Cu | 789.6 | 0.7896 |  | 20240 |  |  |  | 0 | 0.050499 | 0.50499 | 5.0499 | 50.499 | 100.998 | [5 empty] | Ni | 20010
Zn | 3760 | 3.76 |  | 100200 |  |  |  | 0 | 0.25 | 2.5 | 25 | 250 | 500 | 200.4 |  |  |  |  | Cu | 20020
Se | 10904 | 10.904 |  | 20130 |  |  |  | 0 | 0.050225 | 0.50225 | 5.0225 | 50.2246 | 100.4491 | [5 empty] | Zn | 99920
Sr | 3534.4 | 3.5344 |  | 99840 |  |  |  | 0 | 0.2491 | 2.491 | 24.9102 | 249.1018 | 498.2036 | [5 empty] | Se | 20000
Ba | 1353.6 | 1.3536 |  | 20130 |  |  |  | 0 | 0.050225 | 0.50225 | 5.0225 | 50.2246 | 100.4491 | [5 empty] | Sr | 99950
Pb | 161.68 | 0.16168 |  | 20090 |  |  |  | 0 | 0.050125 | 0.50125 | 5.0125 | 50.1248 | 100.2495 | [5 empty] | Ba | 19990
[19 empty] | Pb | 19890
| L7: Intermediate dilution (Only for those in green) | [12 empty] | L6 | [12 empty] | L5 | [11 empty] | L4 | [11 empty] | L3 | [11 empty] | L2
| Req. DF | Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | Dens. (g/mL) | Real DF | Real L6 (ppb) |  | Req. DF | Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | Dens. (g/mL) | Real DF | Real L6 (ppb) |  | Req. DF | Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | Dens. (g/mL) | Real DF | Real L5 (ppb) | Req. DF | Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | Dens. (g/mL) | Real DF | Real L4 (ppb) | Req. DF | Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | Dens. (g/mL) | Real DF | Real L3 (ppb) | Req. DF | Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | Dens. (g/mL) | Real DF | Real L2 (ppb)
Li | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05457 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 186.2736 | 536.0931 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 267.3137 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 13.4935 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 1.3597 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.13719
Na | 50 | 10 | 0.2 | 200 | 0.21945 | 9.8 | 9.8964 | 10.1158 | 10.197 | 0.99204 | 50.6947 | 197239.6994 |  | 40 | 10 | 0.25 | 250 | 0.25364 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 39.0926 | 5045.4479 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 2515.8268 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 126.9947 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 12.7971 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 1.2912
Mg | 50 | 10 | 0.2 | 200 | 0.21945 | 9.8 | 9.8964 | 10.1158 | 10.197 | 0.99204 | 50.6947 | 98609.9867 |  | 40 | 10 | 0.25 | 250 | 0.25364 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 39.0926 | 2522.4716 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 1257.7876 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 63.491 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 6.3979 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.64554
Al | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05457 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 186.2736 | 536.4689 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 267.5011 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 13.503 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 1.3607 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.13729
Ca | 50 | 10 | 0.2 | 200 | 0.21945 | 9.8 | 9.8964 | 10.1158 | 10.197 | 0.99204 | 50.6947 | 197259.4253 |  | 40 | 10 | 0.25 | 250 | 0.25364 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 39.0926 | 5045.9525 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 2516.0784 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 127.0074 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 12.7984 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 1.2913
Cr | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05457 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 186.2736 | 108.0132 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 53.8589 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 2.7187 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 0.27396 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.027642
Mn | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05457 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 186.2736 | 108.0668 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 53.8857 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 2.7201 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 0.2741 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.027656
Fe | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05457 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 186.2736 | 536.63 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 267.5814 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 13.5071 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 1.3611 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.13733
Ni | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05457 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 186.2736 | 107.3689 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 53.5377 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 2.7025 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 0.27233 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.027477
Cu | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05457 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 186.2736 | 108.6574 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 54.1801 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 2.7349 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 0.2756 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.027807
Zn | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05457 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 186.2736 | 537.9184 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 268.2239 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 13.5395 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 1.3644 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.13766
Se | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05457 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 186.2736 | 108.0668 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 53.8857 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 2.7201 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 0.2741 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.027656
Sr | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05457 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 186.2736 | 535.9857 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 267.2602 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 13.4908 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 1.3595 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.13717
Ba | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05457 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 186.2736 | 108.0668 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 53.8857 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 2.7201 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 0.2741 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.027656
Pb | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05457 | 9.7 | 9.8019 | 10.1101 | 9.995 | 1.0115 | 186.2736 | 107.8521 |  | 2 | 10 | 5 | 5000 | 5.0537 | 5 | 5.0839 | 10.1376 | 10.0198 | 1.0118 | 2.0055 | 53.7786 | 10 | 10 | 1 | 1000 | 1.0057 | 9 | 9.0475 | 10.0531 | 9.9342 | 1.012 | 19.8105 | 2.7147 | 10 | 10 | 1 | 1000 | 1.0121 | 9 | 9.0319 | 10.044 | 9.9249 | 1.012 | 9.9237 | 0.27355 | 10 | 10 | 1 | 1000 | 1.015 | 9 | 9.0449 | 10.06 | 9.9407 | 1.012 | 9.911 | 0.027601
| Q7: Intermediate dilution (Only for those in green) | [12 empty] | Q6 | [12 empty] | Q5 | [11 empty] | Q4
| Req. DF | Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | Dens. (g/mL) | Real DF | Real Q7 (ppb) |  | Req. DF | Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | Dens. (g/mL) | Real DF | Real Q6 (ppb) |  | Req. DF | Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | Dens. (g/mL) | Real DF | Real Q5 (ppb) | Req. DF | Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | Dens. (g/mL) | Real DF | Real Q4 (ppb)
Li | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05356 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 189.1354 | 189.1354 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 94.5934 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 4.8134
Na | 50 | 10 | 0.2 | 200 | 0.22044 | 9.8 | 9.8653 | 10.0857 | 10.1701 | 0.99171 | 49.8729 | 200509.5966 |  | 40 | 10 | 0.25 | 250 | 0.25508 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 39.0926 | 255802.8579 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 127936.113 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 6510.0848
Mg | 50 | 10 | 0.2 | 200 | 0.21945 | 9.8 | 9.8964 | 10.1158 | 10.1701 | 0.99171 | 49.8729 | 100254.7983 |  | 40 | 10 | 0.25 | 250 | 0.25508 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 38.9171 | 128478.2139 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 64256.5272 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 3269.7214
Al | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05356 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 189.1354 | 527.8229 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 263.983 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 13.4329
Ca | 50 | 10 | 0.2 | 200 | 0.21945 | 9.8 | 9.8964 | 10.1158 | 10.197 | 0.99171 | 49.8729 | 200509.5966 |  | 40 | 10 | 0.25 | 250 | 0.25508 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 39.0926 | 255802.8579 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 127936.113 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 6510.0848
Cr | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05356 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 189.1354 | 105.7443 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 52.8865 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 2.6912
Mn | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05356 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 189.1354 | 105.7443 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 52.8865 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 2.6912
Fe | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05356 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 189.1354 | 528.616 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 264.3797 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 13.4531
Ni | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05356 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 189.1354 | 105.7972 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 52.913 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 2.6925
Cu | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05356 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 189.1354 | 105.8501 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 52.9394 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 2.6938
Zn | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05356 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 189.1354 | 528.2987 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 264.221 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 13.445
Se | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05356 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 189.1354 | 105.7443 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 52.8865 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 2.6912
Sr | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05356 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 189.1354 | 528.4573 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 264.3003 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 13.449
Ba | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05356 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 189.1354 | 105.6915 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 52.8601 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 2.6898
Pb | [13 empty] | 200 | 10 | 0.05 | 50 | 0.05356 | 9.7 | 9.8163 | 10.125 | 10.01 | 1.0115 | 189.1354 | 105.1627 |  | 2 | 10 | 5 | 5000 | 5.067 | 5 | 5.0668 | 10.1337 | 10.0161 | 1.0117 | 1.9995 | 52.5956 | 10 | 10 | 1 | 1000 | 1.0149 | 9 | 9.0797 | 10.0946 | 9.9751 | 1.012 | 19.652 | 2.6764
Real Concentrations
Analyte | L1 | L2 | L3 | L4 | L5 | L6 | Q4 |  | Li / 7 [#3] | 4.8134 |  | 4.8134
Li | 0 | 0.13719 | 1.3597 | 13.4935 | 267.3137 | 536.0931 | 4.8134 |  | Na / 23 [#3] | 6510.0848 |  | 6510.0848
Na | 0 | 1.2912 | 12.7971 | 126.9947 | 2515.8268 | 5045.4479 | 6510.0848 |  | Mg / 24 [#3] | 3269.7214 |  | 3269.7214
Mg | 0 | 0.64554 | 6.3979 | 63.491 | 1257.7876 | 2522.4716 | 3269.7214 |  | Mg / 26 [#3] | 3269.7214 |  | 3269.7214
Mg | 0 | 0.64554 | 6.3979 | 63.491 | 1257.7876 | 2522.4716 | 3269.7214 |  | Al / 27 [#3] | 13.4329 |  | 13.4329
Al | 0 | 0.13729 | 1.3607 | 13.503 | 267.5011 | 536.4689 | 13.4329 |  | Ca / 40 [#2] | 6510.0848 |  | 6510.0848
Ca | 0 | 1.2913 | 12.7984 | 127.0074 | 2516.0784 | 5045.9525 | 6510.0848 |  | Ca / 43 [#3] | 6510.0848 |  | 6510.0848
Ca | 0 | 1.2913 | 12.7984 | 127.0074 | 2516.0784 | 5045.9525 | 6510.0848 |  | Ca / 44 [#3] | 6510.0848 |  | 6510.0848
Ca | 0 | 1.2913 | 12.7984 | 127.0074 | 2516.0784 | 5045.9525 | 6510.0848 |  | Cr / 52 [#1] | 2.6912 |  | 2.6912
Cr | 0 | 0.027642 | 0.27396 | 2.7187 | 53.8589 | 108.0132 | 2.6912 |  | Cr / 53 [#1] | 2.6912 |  | 2.6912
Cr | 0 | 0.027642 | 0.27396 | 2.7187 | 53.8589 | 108.0132 | 2.6912 |  | Mn / 55 [#3] | 2.6912 |  | 2.6912
Mn | 0 | 0.027656 | 0.2741 | 2.7201 | 53.8857 | 108.0668 | 2.6912 |  | Fe / 56 [#1] | 13.4531 |  | 13.4531
Fe | 0 | 0.13733 | 1.3611 | 13.5071 | 267.5814 | 536.63 | 13.4531 |  | Fe / 56 [#2] | 13.4531 |  | 13.4531
Fe | 0 | 0.13733 | 1.3611 | 13.5071 | 267.5814 | 536.63 | 13.4531 |  | Ni / 60 [#3] | 2.6925 |  | 2.6925
Ni | 0 | 0.027477 | 0.27233 | 2.7025 | 53.5377 | 107.3689 | 2.6925 |  | Cu / 63 [#3] | 2.6938 |  | 2.6938
Cu | 0 | 0.027807 | 0.2756 | 2.7349 | 54.1801 | 108.6574 | 2.6938 |  | Cu / 65 [#3] | 2.6938 |  | 2.6938
Cu | 0 | 0.027807 | 0.2756 | 2.7349 | 54.1801 | 108.6574 | 2.6938 |  | Zn / 66 [#3] | 13.445 |  | 13.445
Zn | 0 | 0.13766 | 1.3644 | 13.5395 | 268.2239 | 537.9184 | 13.445 |  | Zn / 68 [#3] | 13.445 |  | 13.445
Zn | 0 | 0.13766 | 1.3644 | 13.5395 | 268.2239 | 537.9184 | 13.445 |  | Se / 78 [#1] | 2.6912 |  | 2.6912
Se | 0 | 0.027656 | 0.2741 | 2.7201 | 53.8857 | 108.0668 | 2.6912 |  | Se / 78 [#2] | 2.6912 |  | 2.6912
Se | 0 | 0.027656 | 0.2741 | 2.7201 | 53.8857 | 108.0668 | 2.6912 |  | Se / 80 [#1] | 2.6912 |  | 2.6912
Se | 0 | 0.027656 | 0.2741 | 2.7201 | 53.8857 | 108.0668 | 2.6912 |  | Se / 80 [#2] | 2.6912 |  | 2.6912
Se | 0 | 0.027656 | 0.2741 | 2.7201 | 53.8857 | 108.0668 | 2.6912 |  | Se / 82 [#3] | 2.6912 |  | 2.6912
Se | 0 | 0.027656 | 0.2741 | 2.7201 | 53.8857 | 108.0668 | 2.6912 |  | Sr / 88 [#3] | 13.449 |  | 13.449
Sr | 0 | 0.13717 | 1.3595 | 13.4908 | 267.2602 | 535.9857 | 13.449 |  | Ba / 137 [#3] | 2.6898 |  | 2.6898
Ba | 0 | 0.027656 | 0.2741 | 2.7201 | 53.8857 | 108.0668 | 2.6898 |  | Ba / 138 [#3] | 2.6898 |  | 2.6898
Ba | 0 | 0.027656 | 0.2741 | 2.7201 | 53.8857 | 108.0668 | 2.6898 |  | Pb / 208 [#3] | 2.6764 |  | 2.6764
Pb | 0 | 0.027601 | 0.27355 | 2.7147 | 53.7786 | 107.8521 | 2.6764
Pb | 0 | 0.027601 | 0.27355 | 2.7147 | 53.7786 | 107.8521 | 2.6764
Pb | 0 | 0.027601 | 0.27355 | 2.7147 | 53.7786 | 107.8521 | 2.6764
NIST 10x
| Li / 7 [#3] | Na / 23 [#3] | Mg / 24 [#3] | Mg / 26 [#3] | Al / 27 [#3] | Ca / 40 [#2] | Ca / 43 [#3] | Ca / 44 [#3] | Cr / 52 [#1] | Cr / 53 [#1] | Mn / 55 [#3] | Fe / 56 [#1] | Fe / 56 [#2] | Ni / 60 [#3] | Cu / 63 [#3] | Cu / 65 [#3] | Zn / 66 [#3] | Zn / 68 [#3] | Se / 78 [#1] | Se / 78 [#2] | Se / 80 [#1] | Se / 80 [#2] | Se / 82 [#3] | Sr / 88 [#3] | Ba / 137 [#3] | Ba / 138 [#3] | Pb / 208 [#3]
| 1.7792 | 2120.74 | 821.8123 | 821.8123 | 14.4996 | 3302.7918 | 3302.7918 | 3302.7918 | 2.086 | 2.086 | 3.9848 | 10.0311 | 10.0311 | 6.3816 | 2.3273 | 2.3273 | 8.0269 | 8.0269 | 1.224 | 1.224 | 1.224 | 1.224 | 1.224 | 33.0381 | 55.6464 | 55.6464 | 0

## Beech Val
|  | L7 - Initial 1mL diluted to 10 mL | [12 empty] | L6
Analyte |  | Req. DF | Desired Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 (mL) | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | New Dens. (g/mL) | Real DF | Real L7 (ppm) |  | Req. DF | Desired Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 (mL) | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | New Dens. (g/mL) | Real DF | Real L6 (ppm)
Al |  | 10 | 5 | 0.5 | 500 | 0.55101 | 4.5 | 4.4949 | 5.0459 | 4.9687 | 1.0155 | 9.427 | 0.75591 |  | 2 | 5 | 2.5 | 2500 | 2.5481 | 2.5 | 2.5249 | 5.073 | 5.004 | 1.0138 | 2.0218 | 0.37387
Ca |  | 10 | 5 | 0.5 | 500 | 0.55101 | 4.5 | 4.4949 | 5.0459 | 4.9687 | 1.0155 | 9.427 | 2.5854 |  | 2 | 5 | 2.5 | 2500 | 2.5481 | 2.5 | 2.5249 | 5.073 | 5.004 | 1.0138 | 2.0218 | 1.2787
Fe |  | 10 | 5 | 0.5 | 500 | 0.55101 | 4.5 | 4.4949 | 5.0459 | 4.9687 | 1.0155 | 9.427 | 3.4317 |  | 2 | 5 | 2.5 | 2500 | 2.5481 | 2.5 | 2.5249 | 5.073 | 5.004 | 1.0138 | 2.0218 | 1.6973
Sr |  | 10 | 5 | 0.5 | 500 | 0.55101 | 4.5 | 4.4949 | 5.0459 | 4.9687 | 1.0155 | 9.427 | 0.016579 |  | 2 | 5 | 2.5 | 2500 | 2.5481 | 2.5 | 2.5249 | 5.073 | 5.004 | 1.0138 | 2.0218 | 0.0082
Ba |  | 10 | 5 | 0.5 | 500 | 0.55101 | 4.5 | 4.4949 | 5.0459 | 4.9687 | 1.0155 | 9.427 | 0.020173 |  | 2 | 5 | 2.5 | 2500 | 2.5481 | 2.5 | 2.5249 | 5.073 | 5.004 | 1.0138 | 2.0218 | 0.0099776
Pb |  | 10 | 5 | 0.5 | 500 | 0.55101 | 4.5 | 4.4949 | 5.0459 | 4.9687 | 1.0155 | 9.427 | 0.0031419 |  | 2 | 5 | 2.5 | 2500 | 2.5481 | 2.5 | 2.5249 | 5.073 | 5.004 | 1.0138 | 2.0218 | 0.001554
|  | L5 | [12 empty] | L4
|  | Req. DF | Desired Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 (mL) | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | New Dens. (g/mL) | Real DF | Real L5 (ppm) |  | Req. DF | Desired Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 (mL) | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | New Dens. (g/mL) | Real DF | Real L4 (ppm)
Al |  | 10 | 5 | 0.5 | 500 | 0.5 | 4.5 | 5.0461 | 5.5461 | 5.4794 | 1.0122 | 11.245 | 0.033248 |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8813 | 1.012 | 10.1216 | 0.0032849
Ca |  | 10 | 5 | 0.5 | 500 | 0.5 | 4.5 | 5.0461 | 5.5461 | 5.4794 | 1.0122 | 11.245 | 0.11372 |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8813 | 1.012 | 10.1216 | 0.011235
Fe |  | 10 | 5 | 0.5 | 500 | 0.5 | 4.5 | 5.0461 | 5.5461 | 5.4794 | 1.0122 | 11.245 | 0.15094 |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8813 | 1.012 | 10.1216 | 0.014913
Sr |  | 10 | 5 | 0.5 | 500 | 0.5 | 4.5 | 5.0461 | 5.5461 | 5.4794 | 1.0122 | 11.245 | 0.00072921 |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8813 | 1.012 | 10.1216 | 7.2045e-05
Ba |  | 10 | 5 | 0.5 | 500 | 0.5 | 4.5 | 5.0461 | 5.5461 | 5.4794 | 1.0122 | 11.245 | 0.00088729 |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8813 | 1.012 | 10.1216 | 8.7663e-05
Pb |  | 10 | 5 | 0.5 | 500 | 0.5 | 4.5 | 5.0461 | 5.5461 | 5.4794 | 1.0122 | 11.245 | 0.00013819 |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8813 | 1.012 | 10.1216 | 1.3653e-05
|  | L3 | [12 empty] | L2
|  | Req. DF | Desired Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 (mL) | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | New Dens. (g/mL) | Real DF | Real L2 (ppm) |  | Req. DF | Desired Tot. V. (mL) | Sample (mL) | Sample (µL) | Sample (g) | Req. HNO3 (mL) | HNO3 (g) | Tot. W (g) | Tot. vol. (mL) | New Dens. (g/mL) | Real DF | Real L2 (ppm)
Al |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8814 | 1.012 | 10.1202 | 0.00032459 |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8814 | 1.012 | 10.12 | 3.2074e-05
Ca |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8814 | 1.012 | 10.1202 | 0.0011102 |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8814 | 1.012 | 10.12 | 0.0001097
Fe |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8814 | 1.012 | 10.1202 | 0.0014736 |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8814 | 1.012 | 10.12 | 0.00014561
Sr |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8814 | 1.012 | 10.1202 | 7.119e-06 |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8814 | 1.012 | 10.12 | 7.0346e-07
Ba |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8814 | 1.012 | 10.1202 | 8.6623e-06 |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8814 | 1.012 | 10.12 | 8.5595e-07
Pb |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8814 | 1.012 | 10.1202 | 1.3491e-06 |  | 10 | 10 | 1 | 1000 | 1 | 9 | 9 | 10 | 9.8814 | 1.012 | 10.12 | 1.3331e-07
Red/box indicates values which need to be entered
Actual concentrations for calibration curve (ppb)
|  | File: | Sample: | Misc Info: | Date/Time: | ALS vial: | Method: | Conc (ppm) >>> | Al | Ca | Fe | Sr | Ba | Pb
|  |  | Std1 | L1 |  |  |  | Conc (ppm) >>> | 0 | 0 | 0 | 0 | 0 | 0
|  |  | Std2 | L2 |  |  |  | Conc (ppm) >>> | 0.00032459 | 0.0011102 | 0.0014736 | 7.119e-06 | 8.6623e-06 | 1.3491e-06
|  |  | Std3 | L3 |  |  |  | Conc (ppm) >>> | 0.0032849 | 0.011235 | 0.014913 | 7.2045e-05 | 8.7663e-05 | 1.3653e-05
|  |  | Std4 | L4 |  |  |  | Conc (ppm) >>> | 0.033248 | 0.11372 | 0.15094 | 0.00072921 | 0.00088729 | 0.00013819
|  |  | Std5 | L5 |  |  |  | Conc (ppm) >>> | 0.37387 | 1.2787 | 1.6973 | 0.0082 | 0.0099776 | 0.001554
|  |  | Std6 | L6 |  |  |  | Conc (ppm) >>> | 0.75591 | 2.5854 | 3.4317 | 0.016579 | 0.020173 | 0.0031419
|  |  | QC1 | QCL4 |  |  |  | Conc (ppm) >>>
Adj. R2 (rho)

## Unknowns
Nitric Acid Den. (g/mL at 20°C) | 1.034
Final Disolving volume (mL) | 1
Miller Nitric Density (g/mL) | 1.012
Diluting to 2 mL (5x)
Sample Name | Sample # | Mass Sample (g) | Before nitric | After nitric | Mass of 3M nitric added | Total Mass (g) | Total Mass (kg) | Sample Mass (mg) | Volume of Nitric (mL) | Density (g/mL) | Conc Sample in Sol (ppm) | Sample # | Desired Dilution Factor | Desired Total Volume (mL) | Sample (mL) | Sample Vol (μL) | Sample Mass (g) | Required Nitric (mL) | Mass Nitric (g) | Total Weight (g) | Total Volume (mL) | New Density | Real DF
1971 | B21-133 | 0.03676 | 16.4999 | 17.5169 | 1.017 | 1.0537 | 0.0010537 | 36.76 | 0.98355 | 1.0714 | 34884.9348 | B21-133 | 10 | 4 | 0.4 | 400 | 0.4 | 3.6 | 3.6 | 4 | 3.9307 | 1.0176 | 10.528
[14 empty] | 4 |  |  |  |  | 3.6
21/17 | B21-134 | 0.02919 | 15.6479 | 16.484 | 0.83611 | 0.8653 | 0.0008653 | 29.19 | 0.80862 | 1.0701 | 33733.9651 | B21-134 | 10 | 4 | 0.4 | 400 | 0.17753 | 3.6 | 3.6681 | 3.8456 | 3.7905 | 1.0145 | 22.8481 | Reqired NO3 balance | 2.9477 | NO3 Added | 3.0158
16/12 | B21-135 | 0.0148 | 16.5475 | 17.6507 | 1.1033 | 1.1181 | 0.0011181 | 14.8 | 1.067 | 1.0479 | 13237.0961 | B21-135 | 10 | 4 | 0.4 | 400 | 0.44167 | 3.6 | 3.5835 | 4.0252 | 3.9625 | 1.0158 | 9.4011
11/07 | B21-136 | 0.00838 | 16.7351 | 17.8468 | 1.1116 | 1.12 | 0.00112 | 8.38 | 1.0751 | 1.0418 | 7482.0092 | B21-136 | 10 | 4 | 0.4 | 400 | 0.43722 | 3.6 | 3.6079 | 4.0451 | 3.9848 | 1.0151 | 9.4948
06/02 | B21-137 | 0.0005 | 15.2398 | 16.3465 | 1.1067 | 1.1072 | 0.0011072 | 0.5 | 1.0703 | 1.0345 | 451.5896 | B21-137 | 10 | 4 | 0.4 | 400 | 0.43759 | 3.6 | 3.6984 | 4.136 | 4.0776 | 1.0143 | 9.6394
01/97 | B21-138 | 0.01681 | 15.2328 | 16.3421 | 1.1092 | 1.126 | 0.001126 | 16.81 | 1.0728 | 1.0497 | 14928.4217 | B21-138 | 10 | 4 | 0.4 | 400 | 0.44169 | 3.6 | 3.63 | 4.0717 | 4.0077 | 1.016 | 9.5243
96/92 | B21-139 | 0.01107 | 15.4633 | 16.5723 | 1.109 | 1.1201 | 0.0011201 | 11.07 | 1.0726 | 1.0443 | 9882.9579 | B21-139 | 10 | 4 | 0.4 | 400 | 0.44209 | 3.6 | 3.6314 | 4.0735 | 4.0117 | 1.0154 | 9.4766
91/87 | B21-140 | 0.02345 | 15.1764 | 16.274 | 1.0976 | 1.1211 | 0.0011211 | 23.45 | 1.0615 | 1.0561 | 20917.8895 | B21-140 | 10 | 4 | 0.4 | 400 | 0.43778 | 3.6 | 3.6307 | 4.0685 | 4.0022 | 1.0166 | 9.6549
86/82 | B21-141 | 0.0329 | 16.4797 | 17.5853 | 1.1057 | 1.1386 | 0.0011386 | 32.9 | 1.0693 | 1.0648 | 28895.6419 | B21-141 | 10 | 4 | 0.4 | 400 | 0.43188 | 3.6 | 3.6384 | 4.0702 | 4.0008 | 1.0173 | 9.8638
81/77 | B21-142 | 0.02714 | 15.275 | 16.3767 | 1.1017 | 1.1288 | 0.0011288 | 27.14 | 1.0655 | 1.0595 | 24042.5928 | B21-142 | 10 | 4 | 0.4 | 400 | 0.43947 | 3.6 | 3.6363 | 4.0758 | 4.008 | 1.0169 | 9.6624
76/72 | B21-143 | 0.01597 | 14.956 | 16.0675 | 1.1115 | 1.1275 | 0.0011275 | 15.97 | 1.075 | 1.0489 | 14164.4567 | B21-143 | 10 | 4 | 0.4 | 400 | 0.41534 | 3.6 | 3.6322 | 4.0475 | 3.9851 | 1.0157 | 10.0635
4
Blank | B21-146 |  | 15.2713 | 16.3746 | 1.1033 | 1.1033 | 0.0011033 | 0 | 1.067 | 1.034 | 0 | B21-146 | 10 | 4 | 0.4 | 400 | 0.44346 | 3.6 | 3.6359 | 4.0794 | 4.0217 | 1.0143 | 9.3772
5
BELOW=OBSOLETE (IGNORE)
Diluting to 2 mL (5x)
Sample Name | Sample # | Mass Sample (g) | Before nitric | After nitric | Mass of 3M nitric added | Total Mass (g) | Total Mass (kg) | Sample Mass (mg) | Volume of Nitric (mL) | Density (g/mL) | Conc Sample in Sol (ppm) |  | Desired Dilution Factor | Total Volume (mL) | Sample (mL) | Sample Vol (μL) | Sample Mass (g) | Required Nitric (mL) | Mass Nitric (g) | Total Weight (g) | Total Volume (mL) | New Density | Real DF
1971 | B21-133 | 0.03676 | 16.4999 | 17.5169 | 1.017 | 1.0537 | 0.0010537 | 36.76 | 0.98355 | 1.0714 | 34884.9348 |  | 5 | 0.98355 | 0.19671 | 196.7099 | 0.2 | 0.78684 | 0.78 | 0.98 | 0.95743 | 1.0236 | 5.1288
21/17 | B21-134 | 0.02919 | 15.6479 | 16.484 | 0.83611 | 0.8653 | 0.0008653 | 29.19 | 0.80862 | 1.0701 | 33733.9651 |  | 10 | 0.80862 | 0.080862 | 80.8617 | 0.17753 | 0.72776 | 0.65227 | 0.8298 | 0.81044 | 1.0239 | 4.8851
16/12 | B21-135 | 0.0148 | 16.5475 | 17.6507 | 1.1033 | 1.1181 | 0.0011181 | 14.8 | 1.067 | 1.0479 | 13237.0961 |  | 300 | 1.067 | 0.0035566 | 3.5566 | 0.2 | 4.9964 | 4.99 | 5.19 | 5.1217 | 1.0133 | 26.8344
11/07 | B21-136 | 0.00838 | 16.7351 | 17.8468 | 1.1116 | 1.12 | 0.00112 | 8.38 | 1.0751 | 1.0418 | 7482.0092 |  | 5 | 1.0751 | 0.21502 | 215.0174 | 0.2 | 0.86007 | 0.78 | 0.98 | 0.96273 | 1.0179 | 5.0148
06/02 | B21-137 | 0.0005 | 15.2398 | 16.3465 | 1.1067 | 1.1072 | 0.0011072 | 0.5 | 1.0703 | 1.0345 | 451.5896 |  | 5 | 1.0703 | 0.21406 | 214.0619 | 0.2 | 0.85625 | 0.78 | 0.98 | 0.96409 | 1.0165 | 4.9866
01/97 | B21-138 | 0.01681 | 15.2328 | 16.3421 | 1.1092 | 1.126 | 0.001126 | 16.81 | 1.0728 | 1.0497 | 14928.4217 |  | 5 | 1.0728 | 0.21455 | 214.5513 | 0.2 | 0.85821 | 0.78 | 0.98 | 0.96129 | 1.0195 | 5.0452
96/92 | B21-139 | 0.01107 | 15.4633 | 16.5723 | 1.109 | 1.1201 | 0.0011201 | 11.07 | 1.0726 | 1.0443 | 9882.9579 |  | 5 | 1.0726 | 0.21451 | 214.5145 | 0.2 | 0.85806 | 0.78 | 0.98 | 0.96226 | 1.0184 | 5.0246
91/87 | B21-140 | 0.02345 | 15.1764 | 16.274 | 1.0976 | 1.1211 | 0.0011211 | 23.45 | 1.0615 | 1.0561 | 20917.8895 |  | 5 | 1.0615 | 0.2123 | 212.3017 | 0.2 | 0.84921 | 0.78 | 0.98 | 0.96013 | 1.0207 | 5.0699
86/82 | B21-141 | 0.0329 | 16.4797 | 17.5853 | 1.1057 | 1.1386 | 0.0011386 | 32.9 | 1.0693 | 1.0648 | 28895.6419 |  | 5 | 1.0693 | 0.21386 | 213.8646 | 0.2 | 0.85546 | 0.78 | 0.98 | 0.95859 | 1.0223 | 5.1034
81/77 | B21-142 | 0.02714 | 15.275 | 16.3767 | 1.1017 | 1.1288 | 0.0011288 | 27.14 | 1.0655 | 1.0595 | 24042.5928 |  | 5 | 1.0655 | 0.21309 | 213.0928 | 0.2 | 0.85237 | 0.78 | 0.98 | 0.95952 | 1.0213 | 5.0829
76/72 | B21-143 | 0.01597 | 14.956 | 16.0675 | 1.1115 | 1.1275 | 0.0011275 | 15.97 | 1.075 | 1.0489 | 14164.4567 |  | 5 | 1.075 | 0.21499 | 214.9903 | 0.2 | 0.85996 | 0.78 | 0.98 | 0.96143 | 1.0193 | 5.042
Blank | B21-146 |  | 15.2713 | 16.3746 | 1.1033 | 1.1033 | 0.0011033 | 0 | 1.067 | 1.034 | 0 |  | 5 | 1.067 | 0.21341 | 213.4062 | 0.2 | 0.85362 | 0.78 | 0.98 | 0.96417 | 1.0164 | 4.9848

## Total DF
Total DF
B21-134 | 632.9333
B21-135 | 677.7661
B21-136 | 1218.1083
B21-137 | 20634.3442
B21-138 | 607.8067
B21-139 | 918.1834
B21-140 | 437.0467
B21-141 | 320.594
B21-142 | 379.3283
B21-143 | 677.3822
B21-146 | 9.3772
L3A | 2109497.4795
L4A | 208118.0681
L5A | 20567.9546
L6A | 1845.5852
L7A | 914.0627
L5B | 19610.1841